## Content
Tool Box
Version: 2026-08
| Conversion | Not all moisture is the same. The reason for this is the reference point: While the amount of water is sometimes related to volume when calculating the moisture content, elsewhere it is related to mass. And even if two parties have agreed on the volumetric or gravimetric moisture content, confusion can still arise. This is because the moisture content can be based on both the total volume/total mass and the dry volume/dry mass. Misunderstandings are inevitable. The “Conversion” spreadsheet allows you to convert the various moisture specifications. The “Conversion” spreadsheet allows you to convert between the different moisture definitions, thus avoiding misunderstandings from the outset.
| Calibration | When measuring, a physical measurement variable, such as time, is recorded and converted into another measurement variable, such as speed. The conversion is performed using a calibration. Calibration therefore represents the translation from the measurement variable to the output variable. The same applies to moisture measurement. Here, the measured variable is time, too. Specifically, it is the time required for an electrical pulse to travel a certain distance. Because this travel time is directly related to the water content in the measurement volume, the moisture can be derived from it. The conversion from time to moisture depends on the material to be measured and the presence of air (bulk density). Consequently, at least one reference measurement is required to calibrate the probe. The calibration curve determined then describes the relationship between the measured variable and the moisture as the output variable. The calibration curve can be determined and optimized using the “Calibration” spreadsheet.
| Soil Density Correction | One important strength of PICO sensors is its capability to measure volumetric moisture for many soil types with with just one calibration. As only variable, the user needs to take the soil density into account, which causes a constant offset to the "Universal Soil" calibration. The "Soil Density Correction" spread sheet helps you to obtain the correct offset for your soil application.
| Transform Material Calibration | A material calibration is used to convert the measured tp value into a moisture value. The output can be either volumetric moisture (soil) or gravimetric moisture (others). Depending on the application, it may be necessary to convert an existing calibration to a different reference, e.g., if soil moisture is to be represented gravimetrically. After specifying the input reference and the desired output reference, the calibration can be converted using the “Transform Material Calibration” spreadsheet.

## Conversion
Conversion
Expand this line for further information
When to use this spreadsheet? Use this spreadsheet to convert volumetric moisture to gravimetric moisture and vice versa. How to use this spreadsheet? Adjust the light blue highlighted fields to suit your needs: (1) First, enter the dry density of the material to be measured. (2) Under “Calculate Moisture based on,” select the reference state for the moisture specification, i.e., whether the water content is based on the dry mass or the wet mass (= total mass). (3) Finally, enter the moisture value in “Moisture [Input]” and select whether the input is volumetric or gravimetric moisture.
| ρ |  |  | Moisture [Input] |  |  | Moisture [Output] | Typical Bulk Densitys (Gravimetric, dry Material)
Material | 1.4 | kg/dm3 |  | Volumetric (Total) |  |  | Gravimetric (Total) | Sand (dry) | 1.6 | kg/dm3
Water | 1 | kg/dm3 |  | 5.02% |  |  | 3.638% | Gravel (2-8 mm) | 1.4 | kg/dm3
β | 1.4 | − | [5 empty] | Gravel (8-16 mm) | 1.35 | kg/dm3
[8 empty] | Gravel (16-32 mm) | 1.3 | kg/dm3
[8 empty] | Sawdust | 0.2 | kg/dm3
[8 empty] | Wood Chips | 0.3 | kg/dm3
[8 empty] | Biomass | 0.3 | kg/dm3
[8 empty] | Wheat (Grain) | 0.75 | kg/dm3
[8 empty] | Corn (Maize) | 0.72 | kg/dm3
Output
Input | %MVOL (T) |  | %MVOL (D) |  | %MGRAV (T) | %MGRAV (D)
%MVOL (T)
%MVOL (D)
%MGRAV (T)
%MGRAV (D)
𝛽𝑀/(1+(𝛽−1)∙𝑀)
𝑀/(1−𝑀)
𝑀/(𝛽(1−𝑀)+𝑀)
𝑀/(𝛽(1−𝑀))
𝑀/(𝛽+𝑀)
𝑀/𝛽
𝛽𝑀/(1−𝑀)
𝑀/(1−𝑀)
𝛽𝑀/(1+𝛽𝑀)
𝛽𝑀
𝑀/(1+𝑀)
𝑀/(1+𝑀)
𝑀
𝑀
𝑀
𝑀

## Calibration
Calibration
Expand this line for further information
When to use this spreadsheet? Use this spreadsheet to determine a calibration function for the reference points you have determined. How to use this spreadsheet? (1) Enter your reference measurements in the table on the left: (1) Each pair of values contains the measured tp value (raw sensor value) and the reference moisture content determined using a reference method (e.g. by drying). (2) In the “use?” column, you can choose whether the reference point should be included in the calculation of the calibration curve. (2) This allows you to exclude points (e. g. for a test) without having to remove the information from the table. (3) The supporting point is an additional reference point that is useful if only one reference measurement is available or if the existing reference measurements show little variation. (4) The table on the right shows the polynomial coefficients of the calibration curve. You can choose between linear regresseion (default) and higher order polynomials (up to degree of 5). (5) The setting for “Extend [tp]” only affects the display and causes the calibration curve to be extended beyond the reference points. (5) This allows you to evaluate whether the curve outside the known reference points is meaningful. What do I do with the polynomial coefficients? The polynomial coefficients can be transferred to an existing or new calibration using software (SONO Config / PICO Config / TRIME Tool) or the IMKO CONNECT app.
# | tp | % | use? | Calibration Curve | m0 | m1 | m2 | m3 | m4 | m5
Support | 60 | -3 | yes | Linear | -8.7995 | 0.096187 | 0 | 0 | 0 | 0
1 | 150 | 5.5 | yes | Extend [tp] | m0: 10 | R2= | 0.99557
2 | 155 | 5.8 | yes
3 | 160 | 7 | yes
4 | 157 | 4 | no
5 |  |  | yes
6 |  |  | yes
7 |  |  | yes
8 |  |  | yes
9 |  |  | yes
10 |  |  | yes
11 |  |  | yes
12 |  |  | yes
13 |  |  | yes
14 |  |  | yes
15 |  |  | yes

## Soil Density Correction
Soil Density Correction
Expand this line for further information
When to use this spreadsheet? Use this spreadsheet to determine the correction offset m0 for the "Universal Soil Calibraiton" in case of other soil density than 1.41 kg/l. How to use this spreadsheet? (1) Enter the actual soil density in the "Soil Density [kg/l]" input field. (2) Read the calculated correction offset? What do I do with the correction offset? Use the correction offset to adjust the "Universal Soil" calibration on your probe. This requieres a connection to the probe by software (PICO Config / TRIME Tool) or the IMKO CONNECT app. Note: Don't replace the existing m0 offset by the calculated value but add or substract the value calculated below to/from the "Universal Soil" offset m0.
Note: The density correction only applies to the "Universal Soil" calibration!
Soil Density [kg/l] | Correction Offset |  | Change Calibration
1.5 | -1.13 |  | Decrease calibration parameter m0 by 1,13
Reference Soil Density for "Universal Soil": | 1.4067 | kg/l

## Transform Material Calibration
Transform Material Calibration
Expand this line for further information
When to use this spreadsheet? Use this spreadsheet to convert a material calibration into another material calibration based on a different reference state, e. g. from volumetric to gravimetric. How to use this spreadsheet? Adjust the light blue highlighted fields to suit your needs: (1) Specify the reference state for the initial material calibration (input) and for the desired output material calibration (output). (2) Enter the coefficients for the initial material calibration, m0 to m5. (3) In the Settings section, define the material bulk density and the desired polynomial complexity of the transformed calibration.
Moisture [Input] |  | Moisture [Output]
Volumetric (Total) |  | Gravimetric (Total)
Calibration Coefficients
Input Calibration |  | Output Calibration
-21.5725 | m0 | 22.3251
0.34195 | m1 | 42.3631
-0.0018245 | m2 | 15.9843
5.2874e-06 | m3 | 41.8627
-7.0181e-09 | m4 | 182.9314
3.6106e-12 | m5 | 163.7976
Settings
Density, Material | 1.4 | kg/dm3
Density, Water | 1 | kg/dm3
Den.-Relation, β | 1.4 | −
Calibration Curve | Polynomial [5]
Plot |  |  | Test: | tp = | 500 | Input Moisture | 28.42% | Output Moisture | 22.33%
Start, % | 0
End, % | 100 | [7 empty] | R2= | 0.99995

## Tools (hidden)
Tools
Conversion
Expand this line Details
Read Data from Spread Sheet
| Moisture | 0.0502
| β | 1.4
| Input | Volumetric (Total) |  | 1
| Output | Gravimetric (Total) |  | 3
| Input-Output |  |  | 13
Calculate all Scenarios
|  | %MVOL (T) | %MVOL (D) | %MGRAV (T) | %MGRAV (D)
| %MVOL (T) | 0.0502 | 0.052853 | 0.036379 | 0.037752
| %MVOL (D) | 0.0478 | 0.0502 | 0.034616 | 0.035857
| %MGRAV (T) | 0.068897 | 0.073995 | 0.0502 | 0.052853
| %MGRAV (D) | 0.065665 | 0.07028 | 0.0478 | 0.0502
Transform to read for SVERWEIS
| 11 | 0.0502
| 12 | 0.052853
| 13 | 0.036379
| 14 | 0.037752
| 21 | 0.0478
| 22 | 0.0502
| 23 | 0.034616
| 24 | 0.035857
| 31 | 0.068897
| 32 | 0.073995
| 33 | 0.0502
| 34 | 0.052853
| 41 | 0.065665
| 42 | 0.07028
| 43 | 0.0478
| 44 | 0.0502
| Selected | 0.036379
Plot
| Conversion based on dry Material |  |  |  |  | Conversion based on wet Material
|  | β = 2,5 | β = 1,0 | β = 0,25 |  |  | β = 2,5 | β = 1,0 | β = 0,25
| %M_VOL | 2.5 | 1 | 0.25 |  | %M_VOL | 2.5 | 1 | 0.25
| 0% | 0% | 0% | 0% |  | 0% | 0% | 0% | 0%
| 5% | 2% | 5% | 20% |  | 5% | 2.062% | 5% | 17.39%
| 10% | 4% | 10% | 40% |  | 10% | 4.255% | 10% | 30.77%
| 15% | 6% | 15% | 60% |  | 15% | 6.593% | 15% | 41.38%
| 20% | 8% | 20% | 80% |  | 20% | 9.091% | 20% | 50%
| 25% | 10% | 25% | 100% |  | 25% | 11.76% | 25% | 57.14%
| 30% | 12% | 30% | 120% |  | 30% | 14.63% | 30% | 63.16%
| 35% | 14% | 35% | 140% |  | 35% | 17.72% | 35% | 68.29%
| 40% | 16% | 40% | 160% |  | 40% | 21.05% | 40% | 72.73%
| 45% | 18% | 45% | 180% |  | 45% | 24.66% | 45% | 76.6%
| 50% | 20% | 50% | 200% |  | 50% | 28.57% | 50% | 80%
| 55% | 22% | 55% | 220% |  | 55% | 32.84% | 55% | 83.02%
| 60% | 24% | 60% | 240% |  | 60% | 37.5% | 60% | 85.71%
| 65% | 26% | 65% | 260% |  | 65% | 42.62% | 65% | 88.14%
| 70% | 28% | 70% | 280% |  | 70% | 48.28% | 70% | 90.32%
| 75% | 30% | 75% | 300% |  | 75% | 54.55% | 75% | 92.31%
| 80% | 32% | 80% | 320% |  | 80% | 61.54% | 80% | 94.12%
| 85% | 34% | 85% | 340% |  | 85% | 69.39% | 85% | 95.77%
| 90% | 36% | 90% | 360% |  | 90% | 78.26% | 90% | 97.3%
| 95% | 38% | 95% | 380% |  | 95% | 88.37% | 95% | 98.7%
| 100% | 40% | 100% | 400% |  | 100% | 100% | 100% | 100%
Calibration
Expand this line Details
Read Data from Spread Sheet
| List of Measuring Points (unsorted) |  |  |  | Filter & Sort
| # | tp | % |  | tp | %
| 0 | 60 | -3 |  | 60 | -3
| 1 | 150 | 5.5 |  | 150 | 5.5
| 2 | 155 | 5.8 |  | 155 | 5.8
| 3 | 160 | 7 |  | 160 | 7
| 4 |  | 0
| 5 |  | 0
| 6 |  | 0
| 7 |  | 0
| 8 |  | 0
| 9 |  | 0
| 10 |  | 0
| 11 |  | 0
| 12 |  | 0
| 13 |  | 0
| 14 |  | 0
| 15 |  | 0
Calculate Trend Line (RGP)
| Selected Degree for Calibration |  | 1
Degree
|  | 5 | 4 | 3 | 2 | 1 | 0
| 0 | 0 | 0 | 0 | 0 | 60 | 1
| 1 | 0 | 0 | 0 | 0 | 150 | 1
| 2 | 0 | 0 | 0 | 0 | 155 | 1
| 3 | 0 | 0 | 0 | 0 | 160 | 1
| 4 | 0 | 0 | 0 | 0 | 0 | 0
| 5 | 0 | 0 | 0 | 0 | 0 | 0
| 6 | 0 | 0 | 0 | 0 | 0 | 0
| 7 | 0 | 0 | 0 | 0 | 0 | 0
| 8 | 0 | 0 | 0 | 0 | 0 | 0
| 9 | 0 | 0 | 0 | 0 | 0 | 0
| 10 | 0 | 0 | 0 | 0 | 0 | 0
| 11 | 0 | 0 | 0 | 0 | 0 | 0
| 12 | 0 | 0 | 0 | 0 | 0 | 0
| 13 | 0 | 0 | 0 | 0 | 0 | 0
| 14 | 0 | 0 | 0 | 0 | 0 | 0
| 15 | 0 | 0 | 0 | 0 | 0 | 0
|  | m0 | m1 | m2 | m3 | m4 | m5
|  | -8.7995 | 0.096187 | 0 | 0 | 0 | 0
Plot
| Span MAX-MIN | m0: 120
tp | bar: m0
| 50 | m0: -3.9902
| 52.4 | m0: -3.7593
| 54.8 | m0: -3.5285
| 57.2 | m0: -3.2976
| 59.6 | m0: -3.0668
| 62 | m0: -2.8359
| 64.4 | m0: -2.6051
| 66.8 | m0: -2.3743
| 69.2 | m0: -2.1434
| 71.6 | m0: -1.9126
| 74 | m0: -1.6817
| 76.4 | m0: -1.4509
| 78.8 | m0: -1.22
| 81.2 | m0: -0.98916
| 83.6 | m0: -0.75831
| 86 | m0: -0.52746
| 88.4 | m0: -0.29661
| 90.8 | m0: -0.065764
| 93.2 | m0: 0.16509
| 95.6 | m0: 0.39593
| 98 | m0: 0.62678
| 100.4 | m0: 0.85763
| 102.8 | m0: 1.0885
| 105.2 | m0: 1.3193
| 107.6 | m0: 1.5502
| 110 | m0: 1.781
| 112.4 | m0: 2.0119
| 114.8 | m0: 2.2427
| 117.2 | m0: 2.4736
| 119.6 | m0: 2.7044
| 122 | m0: 2.9353
| 124.4 | m0: 3.1661
| 126.8 | m0: 3.397
| 129.2 | m0: 3.6278
| 131.6 | m0: 3.8587
| 134 | m0: 4.0895
| 136.4 | m0: 4.3204
| 138.8 | m0: 4.5512
| 141.2 | m0: 4.7821
| 143.6 | m0: 5.0129
| 146 | m0: 5.2438
| 148.4 | m0: 5.4746
| 150.8 | m0: 5.7055
| 153.2 | m0: 5.9363
| 155.6 | m0: 6.1672
| 158 | m0: 6.398
| 160.4 | m0: 6.6289
| 162.8 | m0: 6.8597
| 165.2 | m0: 7.0905
| 167.6 | m0: 7.3214
| 170 | m0: 7.5522
Calculate R2: Compare actual measurement with calculated measurement
Calculated with Regression
| tp | % |  | y_Avg | Numerator | Denominator
| 60 | -3.0283 |  | 3.825 | 46.968 | 46.5806
| 150 | m0: 5.6285; m3: 3.2526; m4: 2.8056
| 155 | m0: 6.1094; m3: 5.2187; m4: 3.9006
| 160 | m0: 6.5904; m3: 7.6473; m4: 10.0806
|  |  |  | Sum | 63.0866 | 63.3675
|  |  |  | R^2 | 0.99557
Transform Base Alignment (INTERNAL)
Expand this line Details
Read Data from Spread Sheet
Calibration Coefficients
|  | Standard | transformed
| m0 | m0: #REF!; m1: #REF!
| m1 | m0: #REF!; m1: #REF!
Plot
| tp | m0: Moist_Input; m1: Moist_Output
| #REF! | m0: #REF!; m1: #REF!
| #REF! | m0: #REF!; m1: #REF!
Transform Calibration
Expand this line Details
Read Data from Spread Sheet
Input Calibration Coefficients
| m0 | m1 | m2 | m3 | m4 | m5
| -21.5725 | 0.34195 | -0.0018245 | 5.2874e-06 | -7.0181e-09 | 3.6106e-12
| β | 1.4
| transform from | Volumetric (Total) |  | 1
| transform to | Gravimetric (Total) |  | 3
| equal to Code | 13 | >> | 3
| Degree (Selected) | Polynomial [5] |  | 5
Plot / RGP
| Input Calibration Coefficients m0: Moisture from 0 to 1000 tp; Input Calibration Coefficients m3: Filtered to 0-100 % | Calculation of all possibilites for SVERWEIS()
| tp | Moisture |  | tp | Moisture |  | tp | CODE > | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16
| 0 | -21.5725 |  | 110 | 0.034651 |  | 110 | 0.00034651 | 0.00034651 | 0.00034663 | 0.00024753 | 0.00024759 | 0.00034639 | 0.00034651 | 0.00024744 | 0.00024751 | 0.00048504 | 0.00048528 | 0.00034651 | 0.00034663 | 0.00048488 | 0.00048511 | 0.00034639 | 0.00034651
| 10 | -18.3302 |  | 120 | 0.96098 |  | 120 | 0.0096098 | 0.0096098 | 0.009703 | 0.006883 | 0.0069307 | 0.0095183 | 0.0096098 | 0.0068173 | 0.0068641 | 0.013402 | 0.013584 | 0.0096098 | 0.009703 | 0.013275 | 0.013454 | 0.0095183 | 0.0096098
| 20 | -15.422 |  | 130 | 1.7942 |  | 130 | 0.017942 | 0.017942 | 0.01827 | 0.012882 | 0.01305 | 0.017626 | 0.017942 | 0.012654 | 0.012816 | 0.02494 | 0.025578 | 0.017942 | 0.01827 | 0.024504 | 0.025119 | 0.017626 | 0.017942
| 30 | -12.8187 |  | 140 | 2.5484 |  | 140 | 0.025484 | 0.025484 | 0.026151 | 0.018337 | 0.018679 | 0.024851 | 0.025484 | 0.017878 | 0.018203 | 0.035318 | 0.036611 | 0.025484 | 0.026151 | 0.034449 | 0.035678 | 0.024851 | 0.025484
| 40 | -10.4927 |  | 150 | 3.2365 |  | 150 | 0.032365 | 0.032365 | 0.033448 | 0.023334 | 0.023891 | 0.031351 | 0.032365 | 0.022596 | 0.023118 | 0.044732 | 0.046827 | 0.032365 | 0.033448 | 0.043347 | 0.045312 | 0.031351 | 0.032365
| 50 | -8.4178 |  | 160 | 3.8704 |  | 160 | 0.038704 | 0.038704 | 0.040263 | 0.027955 | 0.028759 | 0.037262 | 0.038704 | 0.026902 | 0.027646 | 0.05336 | 0.056368 | 0.038704 | 0.040263 | 0.051401 | 0.054186 | 0.037262 | 0.038704
| 60 | -6.5694 |  | 170 | 4.4609 |  | 170 | 0.044609 | 0.044609 | 0.046692 | 0.032275 | 0.033352 | 0.042704 | 0.044609 | 0.03088 | 0.031864 | 0.061358 | 0.065369 | 0.044609 | 0.046692 | 0.058782 | 0.062453 | 0.042704 | 0.044609
| 70 | -4.9244 |  | 180 | 5.0179 |  | 180 | 0.050179 | 0.050179 | 0.052829 | 0.036363 | 0.037735 | 0.047781 | 0.050179 | 0.034602 | 0.035842 | 0.068868 | 0.073961 | 0.050179 | 0.052829 | 0.065639 | 0.07025 | 0.047781 | 0.050179
| 80 | -3.4612 |  | 190 | 5.5501 |  | 190 | 0.055501 | 0.055501 | 0.058762 | 0.040282 | 0.041973 | 0.052583 | 0.055501 | 0.038132 | 0.039644 | 0.076014 | 0.082267 | 0.055501 | 0.058762 | 0.072099 | 0.077701 | 0.052583 | 0.055501
| 90 | -2.1594 |  | 200 | 6.0657 |  | 200 | 0.060657 | 0.060657 | 0.064573 | 0.04409 | 0.046124 | 0.057188 | 0.060657 | 0.041527 | 0.043326 | 0.082908 | 0.090403 | 0.060657 | 0.064573 | 0.078272 | 0.084919 | 0.057188 | 0.060657
| 100 | -0.99999 |  | 210 | 6.5716 |  | 210 | 0.065716 | 0.065716 | 0.070339 | 0.047839 | 0.050242 | 0.061664 | 0.065716 | 0.044836 | 0.04694 | 0.089647 | 0.098474 | 0.065716 | 0.070339 | 0.084252 | 0.092003 | 0.061664 | 0.065716
| 110 | 0.034651 |  | 220 | 7.0744 |  | 220 | 0.070744 | 0.070744 | 0.076129 | 0.051574 | 0.054378 | 0.06607 | 0.070744 | 0.048101 | 0.050531 | 0.096316 | 0.10658 | 0.070744 | 0.076129 | 0.090116 | 0.099041 | 0.06607 | 0.070744
| 120 | 0.96098 |  | 230 | 7.5794 |  | 230 | 0.075794 | 0.075794 | 0.082009 | 0.055337 | 0.058578 | 0.070454 | 0.075794 | 0.051358 | 0.054138 | 0.10299 | 0.11481 | 0.075794 | 0.082009 | 0.095932 | 0.10611 | 0.070454 | 0.075794
| 130 | 1.7942 |  | 240 | 8.0914 |  | 240 | 0.080914 | 0.080914 | 0.088038 | 0.059164 | 0.062884 | 0.074857 | 0.080914 | 0.054638 | 0.057796 | 0.10973 | 0.12325 | 0.080914 | 0.088038 | 0.10175 | 0.11328 | 0.074857 | 0.080914
| 140 | 2.5484 |  | 250 | 8.6147 |  | 250 | 0.086147 | 0.086147 | 0.094267 | 0.063086 | 0.067334 | 0.079314 | 0.086147 | 0.057966 | 0.061533 | 0.11659 | 0.13197 | 0.086147 | 0.094267 | 0.10763 | 0.12061 | 0.079314 | 0.086147
| 150 | 3.2365 |  | 260 | 9.1525 |  | 260 | 0.091525 | 0.091525 | 0.10075 | 0.067131 | 0.071962 | 0.083851 | 0.091525 | 0.061364 | 0.065375 | 0.12361 | 0.14104 | 0.091525 | 0.10075 | 0.11358 | 0.12814 | 0.083851 | 0.091525
| 160 | 3.8704 |  | 270 | 9.7079 |  | 270 | 0.097079 | 0.097079 | 0.10752 | 0.071321 | 0.076798 | 0.088489 | 0.097079 | 0.064846 | 0.069342 | 0.13083 | 0.15052 | 0.097079 | 0.10752 | 0.11965 | 0.13591 | 0.088489 | 0.097079
| 170 | 4.4609 |  | 280 | 10.2831 |  | 280 | 0.10283 | 0.10283 | 0.11462 | 0.075674 | 0.08187 | 0.093243 | 0.10283 | 0.068425 | 0.073451 | 0.13828 | 0.16046 | 0.10283 | 0.11462 | 0.12585 | 0.14396 | 0.093243 | 0.10283
| 180 | 5.0179 |  | 290 | 10.88 |  | 290 | 0.1088 | 0.1088 | 0.12208 | 0.080207 | 0.087202 | 0.098124 | 0.1088 | 0.07211 | 0.077714 | 0.14597 | 0.17091 | 0.1088 | 0.12208 | 0.13219 | 0.15232 | 0.098124 | 0.1088
| 190 | 5.5501 |  | 300 | 11.4997 |  | 300 | 0.115 | 0.115 | 0.12994 | 0.084931 | 0.092814 | 0.10314 | 0.115 | 0.075906 | 0.082141 | 0.15392 | 0.18192 | 0.115 | 0.12994 | 0.13867 | 0.161 | 0.10314 | 0.115
| 200 | 6.0657 |  | 310 | 12.1433 |  | 310 | 0.12143 | 0.12143 | 0.13822 | 0.089855 | 0.098727 | 0.10828 | 0.12143 | 0.079815 | 0.086738 | 0.16213 | 0.1935 | 0.12143 | 0.13822 | 0.1453 | 0.17001 | 0.10828 | 0.12143
| 210 | 6.5716 |  | 320 | 12.8112 |  | 320 | 0.12811 | 0.12811 | 0.14694 | 0.094986 | 0.10495 | 0.11356 | 0.12811 | 0.083837 | 0.091509 | 0.17061 | 0.20571 | 0.12811 | 0.14694 | 0.15208 | 0.17936 | 0.11356 | 0.12811
| 220 | 7.0744 |  | 330 | 13.5037 |  | 330 | 0.13504 | 0.13504 | 0.15612 | 0.10033 | 0.11151 | 0.11897 | 0.13504 | 0.08797 | 0.096455 | 0.17936 | 0.21857 | 0.13504 | 0.15612 | 0.15899 | 0.18905 | 0.11897 | 0.13504
| 230 | 7.5794 |  | 340 | 14.2205 |  | 340 | 0.1422 | 0.1422 | 0.16578 | 0.10588 | 0.11841 | 0.1245 | 0.1422 | 0.092209 | 0.10157 | 0.18837 | 0.23209 | 0.1422 | 0.16578 | 0.16603 | 0.19909 | 0.1245 | 0.1422
| 240 | 8.0914 |  | 350 | 14.9613 |  | 350 | 0.14961 | 0.14961 | 0.17593 | 0.11164 | 0.12567 | 0.13014 | 0.14961 | 0.096548 | 0.10687 | 0.19763 | 0.24631 | 0.14961 | 0.17593 | 0.17318 | 0.20946 | 0.13014 | 0.14961
| 250 | 8.6147 |  | 360 | 15.7255 |  | 360 | 0.15725 | 0.15725 | 0.1866 | 0.11761 | 0.13328 | 0.13589 | 0.15725 | 0.10098 | 0.11232 | 0.20713 | 0.26124 | 0.15725 | 0.1866 | 0.18043 | 0.22016 | 0.13589 | 0.15725
| 260 | 9.1525 |  | 370 | 16.5123 |  | 370 | 0.16512 | 0.16512 | 0.19778 | 0.12379 | 0.14127 | 0.14172 | 0.16512 | 0.1055 | 0.11795 | 0.21685 | 0.27689 | 0.16512 | 0.19778 | 0.18777 | 0.23117 | 0.14172 | 0.16512
| 270 | 9.7079 |  | 380 | 17.3209 |  | 380 | 0.17321 | 0.17321 | 0.2095 | 0.13016 | 0.14964 | 0.14764 | 0.17321 | 0.1101 | 0.12372 | 0.22678 | 0.29329 | 0.17321 | 0.2095 | 0.19517 | 0.24249 | 0.14764 | 0.17321
| 280 | 10.2831 |  | 390 | 18.1503 |  | 390 | 0.1815 | 0.1815 | 0.22175 | 0.13674 | 0.15839 | 0.15362 | 0.1815 | 0.11477 | 0.12964 | 0.2369 | 0.31045 | 0.1815 | 0.22175 | 0.20262 | 0.2541 | 0.15362 | 0.1815
| 290 | 10.88 |  | 400 | 18.9994 |  | 400 | 0.18999 | 0.18999 | 0.23456 | 0.1435 | 0.16754 | 0.15966 | 0.18999 | 0.11949 | 0.13571 | 0.24721 | 0.32838 | 0.18999 | 0.23456 | 0.21011 | 0.26599 | 0.15966 | 0.18999
| 300 | 11.4997 |  | 410 | 19.8674 |  | 410 | 0.19867 | 0.19867 | 0.24793 | 0.15045 | 0.17709 | 0.16574 | 0.19867 | 0.12427 | 0.14191 | 0.25767 | 0.3471 | 0.19867 | 0.24793 | 0.21761 | 0.27814 | 0.16574 | 0.19867
| 310 | 12.1433 |  | 420 | 20.7531 |  | 420 | 0.20753 | 0.20753 | 0.26188 | 0.15758 | 0.18706 | 0.17186 | 0.20753 | 0.1291 | 0.14824 | 0.26827 | 0.36663 | 0.20753 | 0.26188 | 0.22513 | 0.29054 | 0.17186 | 0.20753
| 320 | 12.8112 |  | 430 | 21.6558 |  | 430 | 0.21656 | 0.21656 | 0.27642 | 0.16489 | 0.19744 | 0.17801 | 0.21656 | 0.13396 | 0.15468 | 0.27901 | 0.38699 | 0.21656 | 0.27642 | 0.23265 | 0.30318 | 0.17801 | 0.21656
| 330 | 13.5037 |  | 440 | 22.5747 |  | 440 | 0.22575 | 0.22575 | 0.29157 | 0.17237 | 0.20826 | 0.18417 | 0.22575 | 0.13886 | 0.16125 | 0.28987 | 0.4082 | 0.22575 | 0.29157 | 0.24015 | 0.31605 | 0.18417 | 0.22575
| 340 | 14.2205 |  | 450 | 23.5094 |  | 450 | 0.23509 | 0.23509 | 0.30735 | 0.18002 | 0.21954 | 0.19034 | 0.23509 | 0.14378 | 0.16792 | 0.30084 | 0.43029 | 0.23509 | 0.30735 | 0.24763 | 0.32913 | 0.19034 | 0.23509
| 350 | 14.9613 |  | 460 | 24.4594 |  | 460 | 0.24459 | 0.24459 | 0.32379 | 0.18784 | 0.23128 | 0.19652 | 0.24459 | 0.14873 | 0.17471 | 0.31191 | 0.45331 | 0.24459 | 0.32379 | 0.25508 | 0.34243 | 0.19652 | 0.24459
| 360 | 15.7255 |  | 470 | 25.4247 |  | 470 | 0.25425 | 0.25425 | 0.34093 | 0.19583 | 0.24352 | 0.20271 | 0.25425 | 0.15369 | 0.1816 | 0.32309 | 0.4773 | 0.25425 | 0.34093 | 0.26251 | 0.35595 | 0.20271 | 0.25425
| 370 | 16.5123 |  | 480 | 26.4055 |  | 480 | 0.26406 | 0.26406 | 0.3588 | 0.204 | 0.25628 | 0.2089 | 0.26406 | 0.15868 | 0.18861 | 0.33436 | 0.50232 | 0.26406 | 0.3588 | 0.2699 | 0.36968 | 0.2089 | 0.26406
| 380 | 17.3209 |  | 490 | 27.4025 |  | 490 | 0.27403 | 0.27403 | 0.37746 | 0.21236 | 0.26961 | 0.21509 | 0.27403 | 0.16369 | 0.19573 | 0.34574 | 0.52844 | 0.27403 | 0.37746 | 0.27727 | 0.38364 | 0.21509 | 0.27403
| 390 | 18.1503 |  | 500 | 28.4167 |  | 500 | 0.28417 | 0.28417 | 0.39697 | 0.22091 | 0.28355 | 0.22129 | 0.28417 | 0.16873 | 0.20298 | 0.35723 | 0.55576 | 0.28417 | 0.39697 | 0.28461 | 0.39783 | 0.22129 | 0.28417
| 400 | 18.9994 |  | 510 | 29.4494 |  | 510 | 0.29449 | 0.29449 | 0.41742 | 0.22968 | 0.29816 | 0.2275 | 0.29449 | 0.17379 | 0.21035 | 0.36884 | 0.58439 | 0.29449 | 0.41742 | 0.29193 | 0.41229 | 0.2275 | 0.29449
| 410 | 19.8674 |  | 520 | 30.5027 |  | 520 | 0.30503 | 0.30503 | 0.43891 | 0.23868 | 0.3135 | 0.23373 | 0.30503 | 0.1789 | 0.21788 | 0.3806 | 0.61447 | 0.30503 | 0.43891 | 0.29925 | 0.42704 | 0.23373 | 0.30503
| 420 | 20.7531 |  | 530 | 31.5789 |  | 530 | 0.31579 | 0.31579 | 0.46154 | 0.24793 | 0.32967 | 0.24 | 0.31579 | 0.18405 | 0.22556 | 0.39252 | 0.64615 | 0.31579 | 0.46154 | 0.30657 | 0.4421 | 0.24 | 0.31579
| 430 | 21.6558 |  | 540 | 32.6811 |  | 540 | 0.32681 | 0.32681 | 0.48547 | 0.25748 | 0.34676 | 0.24631 | 0.32681 | 0.18926 | 0.23344 | 0.40464 | 0.67965 | 0.32681 | 0.48547 | 0.31391 | 0.45754 | 0.24631 | 0.32681
| 440 | 22.5747 |  | 550 | 33.8129 |  | 550 | 0.33813 | 0.33813 | 0.51087 | 0.26735 | 0.36491 | 0.25269 | 0.33813 | 0.19454 | 0.24152 | 0.41698 | 0.71522 | 0.33813 | 0.51087 | 0.32129 | 0.47338 | 0.25269 | 0.33813
| 450 | 23.5094 |  | 560 | 34.9785 |  | 560 | 0.34979 | 0.34979 | 0.53795 | 0.27759 | 0.38425 | 0.25914 | 0.34979 | 0.1999 | 0.24985 | 0.42959 | 0.75313 | 0.34979 | 0.53795 | 0.32872 | 0.4897 | 0.25914 | 0.34979
| 460 | 24.4594 |  | 570 | 36.183 |  | 570 | 0.36183 | 0.36183 | 0.56698 | 0.28825 | 0.40499 | 0.26569 | 0.36183 | 0.20537 | 0.25845 | 0.44252 | 0.79377 | 0.36183 | 0.56698 | 0.33624 | 0.50656 | 0.26569 | 0.36183
| 470 | 25.4247 |  | 580 | 37.432 |  | 580 | 0.37432 | 0.37432 | 0.59826 | 0.29939 | 0.42733 | 0.27237 | 0.37432 | 0.21097 | 0.26737 | 0.4558 | 0.83757 | 0.37432 | 0.59826 | 0.34385 | 0.52405 | 0.27237 | 0.37432
| 480 | 26.4055 |  | 590 | 38.7321 |  | 590 | 0.38732 | 0.38732 | 0.63218 | 0.31108 | 0.45156 | 0.27919 | 0.38732 | 0.2167 | 0.27666 | 0.46951 | 0.88505 | 0.38732 | 0.63218 | 0.3516 | 0.54225 | 0.27919 | 0.38732
| 490 | 27.4025 |  | 600 | 40.0907 |  | 600 | 0.40091 | 0.40091 | 0.66919 | 0.32341 | 0.47799 | 0.28618 | 0.40091 | 0.22261 | 0.28636 | 0.4837 | 0.93686 | 0.40091 | 0.66919 | 0.3595 | 0.56127 | 0.28618 | 0.40091
| 500 | 28.4167 |  | 610 | 41.5158 |  | 610 | 0.41516 | 0.41516 | 0.70986 | 0.33645 | 0.50704 | 0.29336 | 0.41516 | 0.22872 | 0.29654 | 0.49845 | 0.99381 | 0.41516 | 0.70986 | 0.36758 | 0.58122 | 0.29336 | 0.41516
| 510 | 29.4494 |  | 620 | 43.0166 |  | 620 | 0.43017 | 0.43017 | 0.7549 | 0.35032 | 0.53921 | 0.30078 | 0.43017 | 0.23504 | 0.30726 | 0.51382 | 1.0569 | 0.43017 | 0.7549 | 0.37587 | 0.60223 | 0.30078 | 0.43017
| 520 | 30.5027 |  | 630 | 44.6033 |  | 630 | 0.44603 | 0.44603 | 0.80516 | 0.36513 | 0.57511 | 0.30845 | 0.44603 | 0.24162 | 0.31859 | 0.5299 | 1.1272 | 0.44603 | 0.80516 | 0.38441 | 0.62445 | 0.30845 | 0.44603
| 530 | 31.5789 |  | 640 | 46.2868 |  | 640 | 0.46287 | 0.46287 | 0.86174 | 0.38101 | 0.61553 | 0.31641 | 0.46287 | 0.24847 | 0.33062 | 0.54678 | 1.2064 | 0.46287 | 0.86174 | 0.39321 | 0.64802 | 0.31641 | 0.46287
| 540 | 32.6811 |  | 650 | 48.0794 |  | 650 | 0.48079 | 0.48079 | 0.92602 | 0.39811 | 0.66144 | 0.32469 | 0.48079 | 0.25563 | 0.34342 | 0.56454 | 1.2964 | 0.48079 | 0.92602 | 0.40231 | 0.67311 | 0.32469 | 0.48079
| 550 | 33.8129 |  | 660 | 49.9942 |  | 660 | 0.49994 | 0.49994 | 0.99977 | 0.41661 | 0.71412 | 0.33331 | 0.49994 | 0.26314 | 0.3571 | 0.58328 | 1.3997 | 0.49994 | 0.99977 | 0.41174 | 0.69992 | 0.33331 | 0.49994
| 560 | 34.9785 |  | 670 | 52.0455 |  | 670 | 0.52046 | 0.52046 | 1.0853 | 0.43669 | 0.77522 | 0.3423 | 0.52046 | 0.27101 | 0.37175 | 0.60309 | 1.5194 | 0.52046 | 1.0853 | 0.42151 | 0.72864 | 0.3423 | 0.52046
| 570 | 36.183 |  | 680 | 54.249 |  | 680 | 0.54249 | 0.54249 | 1.1857 | 0.45857 | 0.84696 | 0.3517 | 0.54249 | 0.27928 | 0.38749 | 0.62407 | 1.66 | 0.54249 | 1.1857 | 0.43165 | 0.75949 | 0.3517 | 0.54249
| 580 | 37.432 |  | 690 | 56.6213 |  | 690 | 0.56621 | 0.56621 | 1.3053 | 0.48249 | 0.93234 | 0.36152 | 0.56621 | 0.28797 | 0.40444 | 0.64632 | 1.8274 | 0.56621 | 1.3053 | 0.44218 | 0.7927 | 0.36152 | 0.56621
| 590 | 38.7321 |  | 700 | 59.1805 |  | 700 | 0.5918 | 0.5918 | 1.4498 | 0.50874 | 1.0356 | 0.37178 | 0.5918 | 0.29712 | 0.42272 | 0.66994 | 2.0297 | 0.5918 | 1.4498 | 0.45311 | 0.82853 | 0.37178 | 0.5918
| 600 | 40.0907 |  | 710 | 61.9458 |  | 710 | 0.61946 | 0.61946 | 1.6278 | 0.53762 | 1.1627 | 0.38251 | 0.61946 | 0.30674 | 0.44247 | 0.69503 | 2.279 | 0.61946 | 1.6278 | 0.46445 | 0.86724 | 0.38251 | 0.61946
| 610 | 41.5158 |  | 720 | 64.9379 |  | 720 | 0.64938 | 0.64938 | 1.8521 | 0.56951 | 1.3229 | 0.39371 | 0.64938 | 0.31687 | 0.46384 | 0.72167 | 2.5929 | 0.64938 | 1.8521 | 0.4762 | 0.90913 | 0.39371 | 0.64938
| 620 | 43.0166 |  | 730 | 68.1788 |  | 730 | 0.68179 | 0.68179 | 2.1426 | 0.60481 | 1.5304 | 0.40539 | 0.68179 | 0.3275 | 0.48699 | 0.74997 | 2.9996 | 0.68179 | 2.1426 | 0.48836 | 0.9545 | 0.40539 | 0.68179
| 630 | 44.6033 |  | 740 | 71.6921 |  | 740 | 0.71692 | 0.71692 | 2.5326 | 0.644 | 1.809 | 0.41756 | 0.71692 | 0.33866 | 0.51209 | 0.78001 | 3.5456 | 0.71692 | 2.5326 | 0.50092 | 1.0037 | 0.41756 | 0.71692
| 640 | 46.2868 |  | 750 | 75.5027 |  | 750 | 0.75503 | 0.75503 | 3.0821 | 0.68765 | 2.2015 | 0.43021 | 0.75503 | 0.35036 | 0.53931 | 0.81185 | 4.3149 | 0.75503 | 3.0821 | 0.51386 | 1.057 | 0.43021 | 0.75503
| 650 | 48.0794 |  | 760 | 79.6371 |  | 760 | 0.79637 | 0.79637 | 3.9109 | 0.73639 | 2.7935 | 0.44332 | 0.79637 | 0.36258 | 0.56884 | 0.84557 | 5.4753 | 0.79637 | 3.9109 | 0.52717 | 1.1149 | 0.44332 | 0.79637
| 660 | 49.9942 |  | 770 | 84.1233 |  | 770 | 0.84123 | 0.84123 | 5.2986 | 0.791 | 3.7847 | 0.45689 | 0.84123 | 0.37534 | 0.60088 | 0.88121 | 7.418 | 0.84123 | 5.2986 | 0.54081 | 1.1777 | 0.45689 | 0.84123
| 670 | 52.0455 |  | 780 | 88.991 |  | 780 | 0.88991 | 0.88991 | 8.0835 | 0.85238 | 5.7739 | 0.47087 | 0.88991 | 0.38862 | 0.63565 | 0.91881 | 11.3169 | 0.88991 | 8.0835 | 0.55474 | 1.2459 | 0.47087 | 0.88991
| 680 | 54.249 |  | 790 | 94.2716 |  | 790 | 0.94272 | 0.94272 | 16.4567 | 0.9216 | 11.7548 | 0.48526 | 0.94272 | 0.4024 | 0.67337 | 0.9584 | 23.0394 | 0.94272 | 16.4567 | 0.56893 | 1.3198 | 0.48526 | 0.94272
| 690 | 56.6213 |  | 800 | 99.9979 |  | 800 | 0.99998 | 0.99998 | 46903.3152 | 0.99997 | 33502.368 | 0.49999 | 0.99998 | 0.41666 | 0.71427 | 0.99998 | 65664.6413 | 0.99998 | 46903.3152 | 0.58333 | 1.4 | 0.49999 | 0.99998
| 700 | 59.1805
| 710 | 61.9458
| 720 | 64.9379
| 730 | 68.1788
| 740 | 71.6921
| 750 | 75.5027
| 760 | 79.6371
| 770 | 84.1233
| 780 | 88.991
| 790 | 94.2716
| 800 | 99.9979
| 810 | 106.2047
| 820 | 112.9287
| 830 | 120.208
| 840 | 128.083
| 850 | 136.5957
| 860 | 145.7904
| 870 | 155.7129
| 880 | 166.4115
| 890 | 177.9364
| 900 | 190.3399
| 910 | 203.6765
| 920 | 218.003
| 930 | 233.3783
| 940 | 249.8637
| 950 | 267.5228
| 960 | 286.4218
| 970 | 306.6289
| 980 | 328.2152
| 990 | 351.254
| 1000 | 375.8215
|  | selected Code | Format for RGP |  | Format for Plot |  |  |  | 2: Aux. Matrix for RGP
| tp | OUTPUT | INPUT | OUTPUT | tp | IN | OUT [Calc] | OUT [Filt] | bar: 2
| 110 | 0.024753 | 0.034651 | 0.024753 | 110 | 0.034651 | -0.43838 | #N/A | 2: -0.78; 3: -0.43838
| 120 | 0.6883 | 0.96098 | 0.6883 | 120 | 0.96098 | 0.48353 | 0.48353 | 2: -0.76; 3: 0.48353
| 130 | 1.2882 | 1.7942 | 1.2882 | 130 | 1.7942 | 1.2737 | 1.2737 | 2: -0.74; 3: 1.2737
| 140 | 1.8337 | 2.5484 | 1.8337 | 140 | 2.5484 | 1.952 | 1.952 | 2: -0.72; 3: 1.952
| 150 | 2.3334 | 3.2365 | 2.3334 | 150 | 3.2365 | 2.5367 | 2.5367 | 2: -0.7; 3: 2.5367
| 160 | 2.7955 | 3.8704 | 2.7955 | 160 | 3.8704 | 3.0445 | 3.0445 | 2: -0.68; 3: 3.0445
| 170 | 3.2275 | 4.4609 | 3.2275 | 170 | 4.4609 | 3.4907 | 3.4907 | 2: -0.66; 3: 3.4907
| 180 | 3.6363 | 5.0179 | 3.6363 | 180 | 5.0179 | 3.889 | 3.889 | 2: -0.64; 3: 3.889
| 190 | 4.0282 | 5.5501 | 4.0282 | 190 | 5.5501 | 4.2518 | 4.2518 | 2: -0.62; 3: 4.2518
| 200 | 4.409 | 6.0657 | 4.409 | 200 | 6.0657 | 4.5902 | 4.5902 | 2: -0.6; 3: 4.5902
| 210 | 4.7839 | 6.5716 | 4.7839 | 210 | 6.5716 | 4.9141 | 4.9141 | 2: -0.58; 3: 4.9141
| 220 | 5.1574 | 7.0744 | 5.1574 | 220 | 7.0744 | 5.2322 | 5.2322 | 2: -0.56; 3: 5.2322
| 230 | 5.5337 | 7.5794 | 5.5337 | 230 | 7.5794 | 5.5519 | 5.5519 | 2: -0.54; 3: 5.5519
| 240 | 5.9164 | 8.0914 | 5.9164 | 240 | 8.0914 | 5.8798 | 5.8798 | 2: -0.52; 3: 5.8798
| 250 | 6.3086 | 8.6147 | 6.3086 | 250 | 8.6147 | 6.2213 | 6.2213 | 2: -0.5; 3: 6.2213
| 260 | 6.7131 | 9.1525 | 6.7131 | 260 | 9.1525 | 6.581 | 6.581 | 2: -0.48; 3: 6.581
| 270 | 7.1321 | 9.7079 | 7.1321 | 270 | 9.7079 | 6.9626 | 6.9626 | 2: -0.46; 3: 6.9626
| 280 | 7.5674 | 10.2831 | 7.5674 | 280 | 10.2831 | 7.369 | 7.369 | 2: -0.44; 3: 7.369
| 290 | 8.0207 | 10.88 | 8.0207 | 290 | 10.88 | 7.8023 | 7.8023 | 2: -0.42; 3: 7.8023
| 300 | 8.4931 | 11.4997 | 8.4931 | 300 | 11.4997 | 8.2639 | 8.2639 | 2: -0.4; 3: 8.2639
| 310 | 8.9855 | 12.1433 | 8.9855 | 310 | 12.1433 | 8.7547 | 8.7547 | 2: -0.38; 3: 8.7547
| 320 | 9.4986 | 12.8112 | 9.4986 | 320 | 12.8112 | 9.2749 | 9.2749 | 2: -0.36; 3: 9.2749
| 330 | 10.0325 | 13.5037 | 10.0325 | 330 | 13.5037 | 9.8244 | 9.8244 | 2: -0.34; 3: 9.8244
| 340 | 10.5877 | 14.2205 | 10.5877 | 340 | 14.2205 | 10.4025 | 10.4025 | 2: -0.32; 3: 10.4025
| 350 | 11.1638 | 14.9613 | 11.1638 | 350 | 14.9613 | 11.0082 | 11.0082 | 2: -0.3; 3: 11.0082
| 360 | 11.7609 | 15.7255 | 11.7609 | 360 | 15.7255 | 11.6401 | 11.6401 | 2: -0.28; 3: 11.6401
| 370 | 12.3785 | 16.5123 | 12.3785 | 370 | 16.5123 | 12.2968 | 12.2968 | 2: -0.26; 3: 12.2968
| 380 | 13.0162 | 17.3209 | 13.0162 | 380 | 17.3209 | 12.9764 | 12.9764 | 2: -0.24; 3: 12.9764
| 390 | 13.6736 | 18.1503 | 13.6736 | 390 | 18.1503 | 13.6772 | 13.6772 | 2: -0.22; 3: 13.6772
| 400 | 14.35 | 18.9994 | 14.35 | 400 | 18.9994 | 14.3972 | 14.3972 | 2: -0.2; 3: 14.3972
| 410 | 15.045 | 19.8674 | 15.045 | 410 | 19.8674 | 15.1346 | 15.1346 | 2: -0.18; 3: 15.1346
| 420 | 15.758 | 20.7531 | 15.758 | 420 | 20.7531 | 15.8874 | 15.8874 | 2: -0.16; 3: 15.8874
| 430 | 16.4886 | 21.6558 | 16.4886 | 430 | 21.6558 | 16.6541 | 16.6541 | 2: -0.14; 3: 16.6541
| 440 | 17.2366 | 22.5747 | 17.2366 | 440 | 22.5747 | 17.4332 | 17.4332 | 2: -0.12; 3: 17.4332
| 450 | 18.0016 | 23.5094 | 18.0016 | 450 | 23.5094 | 18.2234 | 18.2234 | 2: -0.1; 3: 18.2234
| 460 | 18.7836 | 24.4594 | 18.7836 | 460 | 24.4594 | 19.0239 | 19.0239 | 2: -0.08; 3: 19.0239
| 470 | 19.583 | 25.4247 | 19.583 | 470 | 25.4247 | 19.834 | 19.834 | 2: -0.06; 3: 19.834
| 480 | 20.4002 | 26.4055 | 20.4002 | 480 | 26.4055 | 20.6539 | 20.6539 | 2: -0.04; 3: 20.6539
| 490 | 21.2358 | 27.4025 | 21.2358 | 490 | 27.4025 | 21.4839 | 21.4839 | 2: -0.02; 3: 21.4839
| 500 | 22.0912 | 28.4167 | 22.0912 | 500 | 28.4167 | 22.3251 | 22.3251 | 2: 0; 3: 22.3251
| 510 | 22.9679 | 29.4494 | 22.9679 | 510 | 29.4494 | 23.1791 | 23.1791 | 2: 0.02; 3: 23.1791
| 520 | 23.8677 | 30.5027 | 23.8677 | 520 | 30.5027 | 24.0484 | 24.0484 | 2: 0.04; 3: 24.0484
| 530 | 24.7934 | 31.5789 | 24.7934 | 530 | 31.5789 | 24.936 | 24.936 | 2: 0.06; 3: 24.936
| 540 | 25.7478 | 32.6811 | 25.7478 | 540 | 32.6811 | 25.8459 | 25.8459 | 2: 0.08; 3: 25.8459
| 550 | 26.7349 | 33.8129 | 26.7349 | 550 | 33.8129 | 26.783 | 26.783 | 2: 0.1; 3: 26.783
| 560 | 27.7588 | 34.9785 | 27.7588 | 560 | 34.9785 | 27.7532 | 27.7532 | 2: 0.12; 3: 27.7532
| 570 | 28.8249 | 36.183 | 28.8249 | 570 | 36.183 | 28.7632 | 28.7632 | 2: 0.14; 3: 28.7632
| 580 | 29.9391 | 37.432 | 29.9391 | 580 | 37.432 | 29.8209 | 29.8209 | 2: 0.16; 3: 29.8209
| 590 | 31.1084 | 38.7321 | 31.1084 | 590 | 38.7321 | 30.9355 | 30.9355 | 2: 0.18; 3: 30.9355
| 600 | 32.3406 | 40.0907 | 32.3406 | 600 | 40.0907 | 32.1171 | 32.1171 | 2: 0.2; 3: 32.1171
| 610 | 33.645 | 41.5158 | 33.645 | 610 | 41.5158 | 33.3773 | 33.3773 | 2: 0.22; 3: 33.3773
| 620 | 35.0317 | 43.0166 | 35.0317 | 620 | 43.0166 | 34.729 | 34.729 | 2: 0.24; 3: 34.729
| 630 | 36.5126 | 44.6033 | 36.5126 | 630 | 44.6033 | 36.1864 | 36.1864 | 2: 0.26; 3: 36.1864
| 640 | 38.1008 | 46.2868 | 38.1008 | 640 | 46.2868 | 37.7652 | 37.7652 | 2: 0.28; 3: 37.7652
| 650 | 39.8113 | 48.0794 | 39.8113 | 650 | 48.0794 | 39.4827 | 39.4827 | 2: 0.3; 3: 39.4827
| 660 | 41.661 | 49.9942 | 41.661 | 660 | 49.9942 | 41.3576 | 41.3576 | 2: 0.32; 3: 41.3576
| 670 | 43.669 | 52.0455 | 43.669 | 670 | 52.0455 | 43.4105 | 43.4105 | 2: 0.34
| 680 | 45.857 | 54.249 | 45.857 | 680 | 54.249 | 45.6635 | 45.6635 | 2: 0.36
| 690 | 48.2493 | 56.6213 | 48.2493 | 690 | 56.6213 | 48.1405 | 48.1405 | 2: 0.38
| 700 | 50.8739 | 59.1805 | 50.8739 | 700 | 59.1805 | 50.8674 | 50.8674 | 2: 0.4
| 710 | 53.7622 | 61.9458 | 53.7622 | 710 | 61.9458 | 53.8717 | 53.8717 | 2: 0.42
| 720 | 56.9506 | 64.9379 | 56.9506 | 720 | 64.9379 | 57.1832 | 57.1832 | 2: 0.44
| 730 | 60.4806 | 68.1788 | 60.4806 | 730 | 68.1788 | 60.8335 | 60.8335 | 2: 0.46
| 740 | 64.4 | 71.6921 | 64.4 | 740 | 71.6921 | 64.8563 | 64.8563 | 2: 0.48
| 750 | 68.7646 | 75.5027 | 68.7646 | 750 | 75.5027 | 69.2875 | 69.2875 | 2: 0.5
| 760 | 73.6391 | 79.6371 | 73.6391 | 760 | 79.6371 | 74.1652 | 74.1652 | 2: 0.52
| 770 | 79.1 | 84.1233 | 79.1 | 770 | 84.1233 | 79.5299 | 79.5299 | 2: 0.54
| 780 | 85.2375 | 88.991 | 85.2375 | 780 | 88.991 | 85.4241 | 85.4241 | 2: 0.56
| 790 | 92.1598 | 94.2716 | 92.1598 | 790 | 94.2716 | 91.8931 | 91.8931 | 2: 0.58
| 800 | 99.997 | 99.9979 | 99.997 | 800 | 99.9979 | 98.9845 | 98.9845 | 2: 0.6
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
|  |  |  | 0 | #N/A | #N/A | #N/A | #N/A | 3: #N/A
RGP
|  | 5 | 4 | 3 | 2 | 1 | 0
| 1 | -0.28872 | 0.37015 | -0.47455 | 0.6084 | -0.78 | 1
| 2 | -0.25355 | 0.33362 | -0.43898 | 0.5776 | -0.76 | 1
| 3 | -0.2219 | 0.29987 | -0.40522 | 0.5476 | -0.74 | 1
| 4 | -0.19349 | 0.26874 | -0.37325 | 0.5184 | -0.72 | 1
| 5 | -0.16807 | 0.2401 | -0.343 | 0.49 | -0.7 | 1
| 6 | -0.14539 | 0.21381 | -0.31443 | 0.4624 | -0.68 | 1
| 7 | -0.12523 | 0.18975 | -0.2875 | 0.4356 | -0.66 | 1
| 8 | -0.10737 | 0.16777 | -0.26214 | 0.4096 | -0.64 | 1
| 9 | -0.091613 | 0.14776 | -0.23833 | 0.3844 | -0.62 | 1
| 10 | -0.07776 | 0.1296 | -0.216 | 0.36 | -0.6 | 1
| 11 | -0.065636 | 0.11316 | -0.19511 | 0.3364 | -0.58 | 1
| 12 | -0.055073 | 0.098345 | -0.17562 | 0.3136 | -0.56 | 1
| 13 | -0.045917 | 0.085031 | -0.15746 | 0.2916 | -0.54 | 1
| 14 | -0.03802 | 0.073116 | -0.14061 | 0.2704 | -0.52 | 1
| 15 | -0.03125 | 0.0625 | -0.125 | 0.25 | -0.5 | 1
| 16 | -0.02548 | 0.053084 | -0.11059 | 0.2304 | -0.48 | 1
| 17 | -0.020596 | 0.044775 | -0.097336 | 0.2116 | -0.46 | 1
| 18 | -0.016492 | 0.037481 | -0.085184 | 0.1936 | -0.44 | 1
| 19 | -0.013069 | 0.031117 | -0.074088 | 0.1764 | -0.42 | 1
| 20 | -0.01024 | 0.0256 | -0.064 | 0.16 | -0.4 | 1
| 21 | -0.0079235 | 0.020851 | -0.054872 | 0.1444 | -0.38 | 1
| 22 | -0.0060466 | 0.016796 | -0.046656 | 0.1296 | -0.36 | 1
| 23 | -0.0045435 | 0.013363 | -0.039304 | 0.1156 | -0.34 | 1
| 24 | -0.0033554 | 0.010486 | -0.032768 | 0.1024 | -0.32 | 1
| 25 | -0.00243 | 0.0081 | -0.027 | 0.09 | -0.3 | 1
| 26 | -0.001721 | 0.0061466 | -0.021952 | 0.0784 | -0.28 | 1
| 27 | -0.0011881 | 0.0045698 | -0.017576 | 0.0676 | -0.26 | 1
| 28 | -0.00079626 | 0.0033178 | -0.013824 | 0.0576 | -0.24 | 1
| 29 | -0.00051536 | 0.0023426 | -0.010648 | 0.0484 | -0.22 | 1
| 30 | -0.00032 | 0.0016 | -0.008 | 0.04 | -0.2 | 1
| 31 | -0.00018896 | 0.0010498 | -0.005832 | 0.0324 | -0.18 | 1
| 32 | -0.00010486 | 0.00065536 | -0.004096 | 0.0256 | -0.16 | 1
| 33 | -5.3782e-05 | 0.00038416 | -0.002744 | 0.0196 | -0.14 | 1
| 34 | -2.4883e-05 | 0.00020736 | -0.001728 | 0.0144 | -0.12 | 1
| 35 | -1e-05 | 0.0001 | -0.001 | 0.01 | -0.1 | 1
| 36 | -3.2768e-06 | 4.096e-05 | -0.000512 | 0.0064 | -0.08 | 1
| 37 | -7.776e-07 | 1.296e-05 | -0.000216 | 0.0036 | -0.06 | 1
| 38 | -1.024e-07 | 2.56e-06 | -6.4e-05 | 0.0016 | -0.04 | 1
| 39 | -3.2e-09 | 1.6e-07 | -8e-06 | 0.0004 | -0.02 | 1
| 40 | 0 | 0 | 0 | 0 | 0 | 1
| 41 | 3.2e-09 | 1.6e-07 | 8e-06 | 0.0004 | 0.02 | 1
| 42 | 1.024e-07 | 2.56e-06 | 6.4e-05 | 0.0016 | 0.04 | 1
| 43 | 7.776e-07 | 1.296e-05 | 0.000216 | 0.0036 | 0.06 | 1
| 44 | 3.2768e-06 | 4.096e-05 | 0.000512 | 0.0064 | 0.08 | 1
| 45 | 1e-05 | 0.0001 | 0.001 | 0.01 | 0.1 | 1
| 46 | 2.4883e-05 | 0.00020736 | 0.001728 | 0.0144 | 0.12 | 1
| 47 | 5.3782e-05 | 0.00038416 | 0.002744 | 0.0196 | 0.14 | 1
| 48 | 0.00010486 | 0.00065536 | 0.004096 | 0.0256 | 0.16 | 1
| 49 | 0.00018896 | 0.0010498 | 0.005832 | 0.0324 | 0.18 | 1
| 50 | 0.00032 | 0.0016 | 0.008 | 0.04 | 0.2 | 1
| 51 | 0.00051536 | 0.0023426 | 0.010648 | 0.0484 | 0.22 | 1
| 52 | 0.00079626 | 0.0033178 | 0.013824 | 0.0576 | 0.24 | 1
| 53 | 0.0011881 | 0.0045698 | 0.017576 | 0.0676 | 0.26 | 1
| 54 | 0.001721 | 0.0061466 | 0.021952 | 0.0784 | 0.28 | 1
| 55 | 0.00243 | 0.0081 | 0.027 | 0.09 | 0.3 | 1
| 56 | 0.0033554 | 0.010486 | 0.032768 | 0.1024 | 0.32 | 1
| 57 | 0.0045435 | 0.013363 | 0.039304 | 0.1156 | 0.34 | 1
| 58 | 0.0060466 | 0.016796 | 0.046656 | 0.1296 | 0.36 | 1
| 59 | 0.0079235 | 0.020851 | 0.054872 | 0.1444 | 0.38 | 1
| 60 | 0.01024 | 0.0256 | 0.064 | 0.16 | 0.4 | 1
| 61 | 0.013069 | 0.031117 | 0.074088 | 0.1764 | 0.42 | 1
| 62 | 0.016492 | 0.037481 | 0.085184 | 0.1936 | 0.44 | 1
| 63 | 0.020596 | 0.044775 | 0.097336 | 0.2116 | 0.46 | 1
| 64 | 0.02548 | 0.053084 | 0.11059 | 0.2304 | 0.48 | 1
| 65 | 0.03125 | 0.0625 | 0.125 | 0.25 | 0.5 | 1
| 66 | 0.03802 | 0.073116 | 0.14061 | 0.2704 | 0.52 | 1
| 67 | 0.045917 | 0.085031 | 0.15746 | 0.2916 | 0.54 | 1
| 68 | 0.055073 | 0.098345 | 0.17562 | 0.3136 | 0.56 | 1
| 69 | 0.065636 | 0.11316 | 0.19511 | 0.3364 | 0.58 | 1
| 70 | 0.07776 | 0.1296 | 0.216 | 0.36 | 0.6 | 1
| 71 | 0 | 0 | 0 | 0 | 0 | 0
| 72 | 0 | 0 | 0 | 0 | 0 | 0
| 73 | 0 | 0 | 0 | 0 | 0 | 0
| 74 | 0 | 0 | 0 | 0 | 0 | 0
| 75 | 0 | 0 | 0 | 0 | 0 | 0
| 76 | 0 | 0 | 0 | 0 | 0 | 0
| 77 | 0 | 0 | 0 | 0 | 0 | 0
| 78 | 0 | 0 | 0 | 0 | 0 | 0
| 79 | 0 | 0 | 0 | 0 | 0 | 0
| 80 | 0 | 0 | 0 | 0 | 0 | 0
| 81 | 0 | 0 | 0 | 0 | 0 | 0
| 82 | 0 | 0 | 0 | 0 | 0 | 0
| 83 | 0 | 0 | 0 | 0 | 0 | 0
| 84 | 0 | 0 | 0 | 0 | 0 | 0
| 85 | 0 | 0 | 0 | 0 | 0 | 0
| 86 | 0 | 0 | 0 | 0 | 0 | 0
| 87 | 0 | 0 | 0 | 0 | 0 | 0
| 88 | 0 | 0 | 0 | 0 | 0 | 0
| 89 | 0 | 0 | 0 | 0 | 0 | 0
| 90 | 0 | 0 | 0 | 0 | 0 | 0
| 91 | 0 | 0 | 0 | 0 | 0 | 0
| 92 | 0 | 0 | 0 | 0 | 0 | 0
| 93 | 0 | 0 | 0 | 0 | 0 | 0
| 94 | 0 | 0 | 0 | 0 | 0 | 0
| 95 | 0 | 0 | 0 | 0 | 0 | 0
| 96 | 0 | 0 | 0 | 0 | 0 | 0
| 97 | 0 | 0 | 0 | 0 | 0 | 0
| 98 | 0 | 0 | 0 | 0 | 0 | 0
| 99 | 0 | 0 | 0 | 0 | 0 | 0
| 100 | 0 | 0 | 0 | 0 | 0 | 0
| m0 | m1 | m2 | m3 | m4 | m5
| 22.3251 | 42.3631 | 15.9843 | 41.8627 | 182.9314 | 163.7976 | 0
| 0.052256 | 0.2633 | 0.82504 | 2.0575 | 2.3705 | 4.1543 | #N/A
| 0.99995 | 0.22241 | #N/A | #N/A | #N/A | #N/A | #N/A
| 291236.411 | 94 | #N/A | #N/A | #N/A | #N/A | #N/A
| 86436.3796 | 4.6497 | #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Test
| tp | 500
| Input | 28.4167
| Output | 22.3251